Here are the steps for checking and resetting the lamp hours from the T-Max Manager Pro:

VIEWING LAMP HOURS:

Press the BED NUMBER you want to check then press ENTER

Press CALL MENU, then 3333, then ENTER

GET PARAMS will appear on display, press ENTER

WHICH PARAM will appear on display, press 6, then ENTER

Display will show current lamp hours.

Pressing the NEXT button will show lamp hours on the next room each time it is pressed.

Exit by pressing the CLEAR button.

CHANGING/ RESETTING LAMP HOURS:

Press the BED NUMBER you want to edit, then press ENTER

Press CALL MENU 3333, then ENTER

Display will show GET PARAMS, press NEXT

Display will show SEND PARAMS, press ENTERR

Display will show WHICH PARAM, press 6, then press ENTER

Display will show SEND WHAT DATA, press 0, then press ENTER to reset hours to zero

Press CLEAR to exit

Press CALL MENU 999 to rescan the T-Max